## Header
CERTIFICATE OF TEST
|  | Operator | Dong-Jun CHOI
Signature
|  | Date | 2023-08-10
TARGET INSTRUMENT :
Model Number | Model Description | Serial Number | Date Tested
BZA60 | 60 V/2 A Battery Impedance Analyzer | R07AS2306001
SIF-II BOARD INFORMATION:
Serial Numder | Mac Address | Board Version | Firmware Version
R07AS2306001 | 00-1B-C5-081907 | 1. 1. 0. 0 | 6. 0. 6. 0
CHANNEL INFORMATION: |  | Channel Count | 1
CHANNEL | Serial Numder | Firmware Version | ZIM Board version
1 | IM600012389 | 4. 1. 0. 0 | 1. 1. 0. 0
EQUIPMENT USED :
Model Number | Model Description | Trace Number | Cal Due date
Keysight Technologies 34470A | Digital Multimeter | MY60039370 | 2016-12-19
Yokogawa CA71 | Handy Multifunction Calibrator
SOFTWARE USED :
Model Number | Model Description
ZM 1. 0. 8. 0 | Smart Manager
ZIVE ZMAN 2.2 f4 | Impedance Spectroscopy Analysis & Presentation
TARGET ACCURACY :
Item | Range | Target Accuracy (1KHz filters are used)
Voltage Measurement | 60 V | PST Setting ±1mV ±0.02% of PST Setting
| 6 V | PST Setting ±1mV ±0.02% of PST Setting
Current Control | 1A ~ 200uA | GAL Setting ±10pA ±0.10% of Range
Temperature Measurement | 1nA ~ 1 A | DMM Reading ±10pA ±0.10% of Range
WonATech Co.,Ltd.

## CH1
S/N #IM600012389
MODE | SET VALUE | DMM READING | ZIM READING | ERROR | PASS
DC Voltage | 6 | 6.011 | 6.0123 | 0.001292 | OK
| 20 | 20.018 | 20.0171 | -0.000928 | OK
| 40 | 40.027 | 40.0278 | 0.000833 | OK
| 60 | 60.042 | 60.043 | 0.001048 | OK
with 60V range ( Gain : 0.99976, Offset : -0.00998 )
DC Voltage | 1 | 0.9998 | 0.99994 | 0.000138 | OK
| 2 | 1.9998 | 1.9997 | -8e-05 | OK
| 4 | 3.9997 | 3.9996 | -5.4e-05 | OK
| 6 | 5.9997 | 5.9998 | 0.000108 | OK
with 6V range ( Gain : 0.99985, Offset : -0.000471 )
Temperature | -50 | -50 | -50.016 | -0.015999 | OK
| 0 | 0 | 0.018007 | 0.018007 | OK
| 50 | 50 | 50.002 | 0.001998 | OK
| 100 | 100 | 100.006 | 0.005995 | OK
with PT100 sensor (Gain : 1.00032, Offset : 0.132078 )
MODE | SET VALUE | Offset(Radian) | DMM READING | ERROR | PASS
DC Current Control | 1000 | 3.249 | 1 | 0 | OK
| 200 | 3.356 | 0.2 | 0 | OK
| 100 | 3.2882 | 0.1 | 0 | OK
| 20 | 3.395 | 0.02 | 0 | OK
| 10 | 3.294 | 0.01 | 0 | OK
| 2 | 3.403 | 0.002 | 0 | OK
| 1 | 3.2985 | 0.001 | 0 | OK
| 0.2 | 3.432 | 0.0002 | 0 | OK
The set value is a fixed value of 50% of each current range.
Freq (Hz) | Zreal (Ω) | Zimag (Ω) | Zmod (Ω) | Zphase (°) | Pass
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 10mohm (9.987261mohm/ 64.476nH) resistor @ 2A range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 10mohm (9.987261mohm/ 64.476nH) resistor @ 400mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 100mohm(99.952mohm/ 112.196nH) @ 200mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 100mohm(99.952mohm/ 112.196nH) @ 40mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 1ohm(0.997961ohm/ 84.638nH) @ 20mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 1ohm(0.997961mohm/ 84.638nH) @ 4mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 10ohm(9.999115 ohm/ 80nH) @ 2mA range.
10 | 0 | 0 | 0 | 0 | OK
1 | 0 | 0 | 0 | 0 | OK
0.1 | 0 | 0 | 0 | 0 | OK
with 10ohm(9.999115 ohm/ 80nH) @ 400uA range.